Ionisation bias undermines the use of MALDI for estimating peptide deamidation: synthetic peptide studies demonstrate ESI gives more reliable response ratios

J.P Simpson1, M. Fascione1, E. Bergström1, 2, J. Wilson1,4, M.J Collins3, K.E.H Penkman1, J Thomas-Oates1,2

1Department of Chemistry, University of York, York

2Centre of Excellence in Mass Spectrometry, University of York, York

3Department of Archaeology, University of York, York

4Department of Mathematics, University of York, York

Abstract

RATIONALE: Although mass spectrometry is routinely used to determine deamination in peptide mixtures, the effects of ionisation source choice have not yet been investigated. In particular, matrix-assisted laser desorption/ionisation (MALDI) has become a popular tool with which to measure levels of glutamine deamidation in ancient proteins. Here we use model synthetic peptides to rigorously compare MALDI and electrospray ionisation.

METHODS:  We use two synthetic peptides, with glutamine (Q) in one substituted for glutamic acid (E) in the other, to investigate the suitability of MALDI and ESI sources for the assessment of deamidation in peptides using mass spectrometry. We also compare measurements of the same Q and E-containing peptide mixtures using two different mass analysers (time of flight and Fourier-transform ion cyclotron resonance).

RESULTS: When standard mixtures of the Q- and E-containing peptides were analysed using MALDI, under-representation of the E-containing peptide was observed. This observation was consistent between analyses carried out using either TOF or FT-ICR-MS. When the same mixtures were analysed using ESI FT-ICR-MS, no ionisation bias was observed.

CONCLUSIONS: MALDI may not be a suitable ionisation method for the determination of deamidation in peptide mixtures. However, ESI was successfully used to determine the ratio in known mixtures of Q and E containing peptides. These preliminary observations warrant further investigation into ionisation bias when measuring deamidation in other peptide sequences.

Keywords: Ionsiation, peptide, deamidation, mass spectrometry

# Introduction

Deamidation of proteins has been described as a molecular clock[1], with its rate dependent on a few known factors; such as the residue in which it occurs, with asparagine (N) found to be less stable than glutamine (Q).[2] Deamidation measurements provide a versatile tool for the investigation of protein stability and diagenetic changes. It is therefore not surprising that deamidation has been used in a broad range of research areas, from modern medicinal applications such as those in biotherapeutics[3], to the investigation of deamidation products in relation to protein aging, such as α-crystallin in eye protein[4,5],Studies on the proteome of human hair[6] shotgun proteomic applications assessing deamidation in complex biological samples[7] antibody and peptide therapies [8,9] as well as more recent applications such as the use of HDX-MS to investigate the effects of deamidation on monoclonal antibody structural confirmations. [10]

In addition to this there have been a wide range of archaeological applications, with proteins recovered from ancient artefacts from sites as old as 3.8 million years in age.[11] This has resulted in a wide range of sample types being reported in the literature, including: keratin in wool[12,13], keratin in mummified skin[14], collagen in bone [15-16], protein binders in paint[17]. Understanding the mechanisms and conditions that drive this reaction is important when trying to measure and quantify levels of deamidation. A number of studies in the literature have used analytical techniques such as liquid chromatography and, more recently, mass spectrometry to measure levels of Q and/or N deamidation in biological[18,19] and synthetic peptides.[20-21] In order to determine the ratios of glutamine to glutamic acid (E) using mass spectrometry, it has generally been assumed by researchers that each of the two forms of the peptide (deamidated and undeamidated) have equal ionisation efficiencies. If this assumption is correct, the ratios of the peak intensities of the Q and E containing species will be directly proportional to the concentration ratios of the two peptide forms.

Matrix-assisted laser desorption/ionisation (MALDI) coupled to a time of flight (TOF) mass spectrometer has been routinely used for the analysis of peptide mixtures and, over recent years, for the estimation of levels of deamidation in various sample types.[22, 23, 12, 13, 16] One disadvantage of this instrumentation is that due to the insufficient resolving power of the mass analyser, it is not possible to fully resolve the signals for the deamidated (+0.984 Da) and amidated peptide forms. For example, the mono-isotopic signal of the deamidated (E-containing) peptide overlaps with the signal for the undeamidated (Q-containing) species. The deamidated mono-isotope differs from the amidated second isotope by only 0.02 Da, with isotopic envelopes becoming more convoluted with increasing deamidation events. In order to calculate the Q/E ratio, a genetic algorithm has been used to deconvolute the overlapping isotopic distributions of the two peptide forms.[24]

Authentic peptide standards are required in order to investigate the accuracy of both the ionisation methods, and the efficiency of the algorithm [24] for the estimation of Q or N deamidation in biological samples. Consequently, in this study two peptides have been synthesised with the amino acid sequences YAYGOGQVG and YAYGOGEVG, where O is used to denote hydroxyproline.  The synthesised peptides were analysed using LC-MS to assess their purity and the two peptide forms were mixed in known ratios.  Analysis was undertaken using MALDI-TOF-MS and MALDI-FT-ICR-MS to determine levels of Q deamidation.  Through this experiment, we therefore assess the applicability of MALDI for the estimation of Q/E ratios, and the efficiency of the previously published deconvolution method.[15,24] In addition, the mixtures were also analysed using the ESI source on the FT-ICR instrument for direct comparison.

# Methods

## Solid phase peptide synthesis

When designing the peptides, several factors were taken into consideration. The peptides needed to be greater than 800 Da in mass, so as not to fall in the same region of the spectrum as the CHCA MALDI matrix peaks. The length of the peptide is also important, as longer peptides have more chance of secondary structures developing during synthesis, as well as generally decreasing peptide purity due to an increase in the number of coupling steps. Therefore, to maximize peptide purity, the peptide was chosen to be no longer than nine amino acids in length. The original context for this study was to investigate the measurement of deamidation in bioarchaeological studies (which commonly look at collagen-derived peptides). However, we acknowledge that the study of deamidation in biological samples is a much wider question. These factors were taken into account when choosing the amino acid composition of the synthetic peptide. Similarities between the peptide presented here, and a typical collagen sequence are, the inclusion of a hydroxyproline residue, and a third of the amino acid composition is glycine (G). As glycine is the smallest amino acid with a residue mass of 57 Da, two tyrosine residues (Y, residue mass of 163 Da) were included in the peptide to offset the small size of the glycines. Tyrosine was also useful as it contains a UV chromophore, enabling the peptides to conveniently be detected and quantified using HPLC with UV detection. The masses of the product peptides, YAYGOGQVG and YAYGOGEVG are 926 Da and 927 Da respectively. The details of the peptide synthesis are provided in the supplementary information (S1).

### Measuring levels of deamidation using mass spectrometry

This study used two mass spectrometers: a Fourier-transform ion cyclotron resonance mass spectrometer (FT-ICR-MS) and a time of flight mass spectrometer (TOF-MS) to generate spectra of the two product peptides (YAYGOGQVG and YAYGOGEVG) combined in various known proportions. These two synthetic peptides are referred to throughout as product peptide Q and product peptide E respectively. The FT-ICR-MS is a high-resolution instrument and was fitted with MALDI and ESI ionisation sources. The second instrument was a lower resolution time of flight instrument, with a fixed MALDI source.  When using the FT-ICR-MS it is possible to resolve the two peptide forms, as the nth peak of the deamidated peptide signal is resolved from the (n+1)th isotopic peak of the undeamidated signal.  The overlapping isotopic distributions in the lower resolution MALDI-TOF-MS data were deconvoluted using the algorithm described in Wilson et al., 2012.[24]

## Peptide purity analysis by HPLC-MS and MALDI-TOF-MS

For analysis using HPLC an aliquot of the lyophilised peptides were re-suspended in water at a concentration of 10 ppm. Each peptide was analysed using an HPLC-HCTultra PTM Discovery System (Bruker Daltonics), fitted with a symmetry C18 3.5 µm (4.6 x 7.5 mm) column (Waters) using mobile phases of acetonitrile (A) and water (B). The elution was isocratic with a flow rate of 1 mL/min and a total run time of 9 minutes.

For analysis by MALDI or ESI-MS the peptides were prepared as follows. 2.02 mg of product peptide Q and 1.96 mg of product peptide E were transferred into separate plastic PP microfuge tubes and dissolved in 2 mL of a 50:50 mixture of ACN and purified water. The resulting 1000 ppm stock solutions were diluted to 10 ppm and mixed in the following weight ratios of Q/E: 100/0, 90/10, 70/30, 50/50, 30/70, 10/90 and 0/100. These mixtures were then used as calibration standards

## Measuring the ratios of the product Q and E peptides using MALDI-TOF, MALDI-FT-ICR and ESI-FT-ICR

Once the purity of each peptide was determined, the set of seven-point calibration standards was then analysed using two different mass spectrometers and two ionisation sources. The two instruments used were the Bruker ultraflex III (TOF) and the Bruker solariX XR 9.4T (FT-ICR). The ultraflex was used with its fixed MALDI source in positive ion mode with 800 laser shots per sample acquisition. The data were acquired using flexcontrol software version 3.0 (Bruker Daltonics). Each spot was analysed in reflector mode using a smartbeam™ Nd:YAG laser (355 nm). Spectra were analysed using flexAnalysis software version 3.0 (Bruker Daltonics). The solariX was operated with either an ESI (direct infusion) or MALDI source in positive ion mode (800 laser shots per sample acquisition) with a smartbeam™ Nd:YAG laser (355 nm). Spectra were acquired using the solariXcontrol software and processed with DataAnalysis version 4.2 (Bruker Daltonics). Each sample was analysed in triplicate and the average values of the peak intensities are reported.

# Results

## Purity analysis of peptides by LC-MS and MALDI-TOF-MS

### Demonstration of peptide sequences

Product ion analysis of the spots prepared with either 100 % product peptide E or product peptide Q was carried out on the MALDI-TOF/TOF instrument. The assignment of the sequence of each of the peptides was made on interpretation of the product ion spectra. The product ion spectrum for product peptide E showed a mixture of both the Q- and E-containing peptides (the precursor selection window on the instrument used is >1 m/z unit). Harsh acidic treatments were used to cleave the synthetic peptides from the resin; it is therefore possible that some of the glutamine in the product peptide Q may have undergone minor deamidation during synthesis.

### Analysis of the purity of product peptides Q and E

To assess the purity of the two synthesised peptides, each product was analysed separately by reversed phase LC-ESI-MS. UV absorbance (210 – 380 nm) data were also collected. As expected, similar absorbance patterns were observed for each of the two peptides. Using this LC-MS approach the product peptides Q and E were resolved chromatographically. Details are shown in the supplementary information (S2). On the basis of the peak areas observed in the UV chromatogram, product peptide Q is estimated to be ~ 95 % pure and to contain 4 % of peptide E, assuming that all components give a similar UV response. From this analysis it is not possible to distinguish whether the small amount of peptide E in product peptide Q arises from deamidation during peptide synthesis, or from the presence of glutamic acid with the glutamine precursor used to make the peptide. Product peptide E was estimated to be ~ 95 % pure and to contain ~3 % of product peptide Q. The structures of product peptides Q and E are shown (Figure 1) with the corresponding amino acid sequences.

Figure 1: The structures of product peptides E and Q. Product peptide E (top) has the amino acid sequence YAYGOGEVG with a mass of 927 Da. Product peptide Q has the amino acid sequence YAYGOGQVG and a mass of 926 Da.

## Comparison of the ratio of product peptide Q to product peptide E signals, measured using different mass spectrometers and ionisation sources.

To investigate possible differences in ionisation efficiency of product peptides, the seven samples of differing Q/E ratios (Table 1) were analysed directly, without chromatographic separation, using both ESI-FT-ICR-MS and MALDI-FT-ICR-MS, making possible direct comparison of the two sets of results. In addition, the seven samples were also analysed using MALDI-TOF-MS. This analysis was carried out in order to compare the estimated ratios of Q/E containing peptides obtained using the deconvolution method described in Wilson et al., 2012 with those results from the FT-ICR experiments in which the different peptides are resolved.

Table 1: corrected percentage of Q and E containing peptides calculated using purity obtained from LC-MS analysis

| Theoretical percentage of Q | Theoretical percentage of E | Corrected percentage of Q taking into account peptide purity | Corrected percentage of E taking into account peptide purity |
| --- | --- | --- | --- |
| 0 | 100 | 3 | 97 |
| 10 | 90 | 12 | 88 |
| 30 | 70 | 31 | 69 |
| 50 | 50 | 50 | 50 |
| 70 | 30 | 69 | 31 |
| 90 | 10 | 87 | 13 |
| 100 | 0 | 96 | 4 |

### FT-ICR-MS Results

The peak intensities of the first three Q- and E-containing peptide peaks in the isotopic distributions were measured using the Bruker solariX XR 9.4T FT-ICR-MS for each mixture using the two different ionisation methods (MALDI and ESI). Each peptide mixture was analysed six times and the average peak intensity of either product peptide Q or E was plotted against their corresponding theoretical percentages (Figure 2).

When the peptide mixtures were analysed using MALDI (Figure A and B), peptide Q consistently produced higher peak intensities than peptide E, regardless of the percentage concentration of the peptide. When using the ESI source, (Figure 2, C and D the percentage of the Q-containing peptide was predicted well across the mixture range, with the average difference between predicted and measured values being ~ 1%. The percentage of peptide Q in the same mixtures determined using MALDI resulted in a greater difference between predicted and measured values, with an average difference of ~ 15%, and differences ranging from ~ 3 - 30%. The errors were greatest within the 10 - 70% Q range, with calculated values for percentage Q overestimated by ~17-30%. It is evident that during MALDI ionisation in positive ion mode, ionisation of the less acidic Q-containing peptide is favoured.

Figure 2: Plots showing measured percentage of peptides Q and E, determined using MALDI (A and B) and ESI (C and D) against the known percentage. When using MALDI there is an underestimation of product peptide E.

### TOF-MS Results

The peptide mixtures were also analysed using MALDI on a TOF mass spectrometer. The percentage of Q-containing peptide in the seven peptide mixtures was estimated using the deconvolution method of Wilson et al. (2012) (Figure 3). The differences between predicted and measured values obtained for the percentage of Q-containing peptide ranged from ~ 3-35 %, with the level of the product peptide E consistently underestimated. The results obtained using MALDI-TOF-MS showed a similar curvilinear relationship to that observed using MALDI on the FT-ICR mass spectrometer. The Wilson et al., (2012) algorithm[24] assumes equal ionisation efficiencies for Q and E-containing peptides and could perhaps be improved by taking ionisation bias into account. However, the study of multiple peptide sequences in a range of biological matrices would be required to investigate a possible correction factor.

Figure 3: Ratios of Q- and E-containing peptides, obtained using MALDI-TOF-MS (B).  Due to an over estimation of Q-containing peptide a curvilinear relationship is observed. The same samples were analysed using a higher resolution FT-ICR-MS (A), with the data showing a similar relationship.

### Investigating effects of peptide concentration, sample matrix and laser power on the obtained measurements of product peptides Q and E

As there appeared to be significant ionisation bias when measuring ratios of Q and E-containing peptides using MALDI, this was explored further. In addition to comparing ionisation sources we also investigated whether there were any effects caused by the concentrations of the peptide mixtures, the MALDI laser intensity, and the effects of the biological matrix. Known ratios of Q and E product peptides were spiked into three different tryptic protein digests: bovine serum albuminm (BSA), cytochrome C (CytC) and collagen extracted from a Roman cow bone (Tanner Row collagen extract).

In order to investigate the effects of concentration, the product peptide mixtures were analysed as standard mixtures in solvent, at three different final concentrations (2.5, 5 and 10 ppm). For each concentration three product peptide mixtures were analysed (30, 50 and 70 percent Q), using either ESI-FT-ICR-MS or MALDI-FT-ICR-MS (Figure 4).

Figure 4: Measured percentage of Q in product peptide mixtures at three concentrations (2.5, 5 and 10 ppm). The data was obtained using either ESI-FT-ICR-MS (A) or MALDI-FT-ICR-MS (B).

When using ESI it was observed that the measured percentage Q across the three peptide ratios was consistent at concentrations of 2.5 and 5 ppm. For the mixture containing 50 % product peptide Q, a measured percentage of 48 was obtained from both the 2.5 and 5 ppm solutions, while from the 10 ppm solution, the measured level was slightly lower at 45 %. For the mixture containing 70 % product peptide Q, the same percentage Q (67 %) was measured from both the 2.5 and 5 ppm solutions, with the 10 ppm solution resulting in a slightly higher measured percentage of 75. When analysing the same mixtures using MALDI, the extent of overestimation of peptide Q slightly increases with the overall concentration of the peptide mixture.

When using MALDI, one of the key instrument parameters that can be tuned is the percentage laser power used. The percentage laser power required to obtain good ionisation of a sample varies with sample type and instrument. As this is a key variable for MALDI analysis we tested the same product peptide mixtures (30, 50 and 70 % Q) at three concentrations (2.5, 5 and 10 ppm), using three different laser powers (25, 30 and 35 %) (Figure 5). We found that although the laser power used affected the quality of the resulting spectra, with a decrease in S/N with decreasing laser power, it does not appear to significantly alter the measured levels of Q and E-containing peptides.

Figure 5: Measured percentage of Q in product peptide mixtures at three concentrations, (A) 2.5 ppm, (B) 5 ppm and (C) 10 ppm analysed using at three laser power percentage (35, 30 and 25 %).

We have shown that there is an overestimation of product peptide Q levels when using MALDI to analyse standard mixtures of two product peptides. However, biological samples such as protein tryptic digests are generally complex mixtures, which are likely to contain a number of Q- and E-containing peptides. In order to test how accurately we can measure Q- and E-product peptides in such biological mixtures, standard peptide mixtures were spiked into tryptic digests of each of three proteins (bovine serum albumin (BSA), cytochrome C (CytC), and collagen extracted from an archaeological cow bone dating from the Roman period. Each of the three digests was analysed spiked with two different concentrations of product peptides (Table 2). Each sample was analysed mass spectrometrically six times and in each spectrum, the first three peaks in the isotopic distribution were summed and used to calculate the level of Q- and E-product peptides. An average of the Q- and E- %s derived from the six replicates was compared with the known percentage of Q for each sample.

Table 2: Measured percentage Q in peptide mixtures spiked into tryptic protein digests (BSA, CytC and Collagen) analysed using either ESI or MALDI.

|  |  |  | ESI |  | MALDI |  |
| --- | --- | --- | --- | --- | --- | --- |
| Biological matrix | [peptide spike] (ppm) | Known % Q | Measured % Q | Standard deviation | Measured % Q | Standard deviation |
| BSA | 5 | 31 | 34 | 2.98 | 58 | 2.21 |
| BSA | 10 | 31 | 34 | 1.63 | 60 | 0.83 |
| BSA | 5 | 50 | 64 | 3.63 | 74 | 0.42 |
| BSA | 10 | 50 | 56 | 0.59 | 74 | 0.42 |
| BSA | 5 | 69 | 85 | 1.41 | 85 | 1.28 |
| BSA | 10 | 69 | 77 | 1.65 | 87 | 0.73 |
| CytC | 5 | 31 | 29 | 0.29 | 56 | 3.00 |
| CytC | 10 | 31 | 31 | 3.78 | 61 | 0.46 |
| CytC | 5 | 50 | 47 | 2.21 | 57 | 1.13 |
| CytC | 10 | 50 | 45 | 0.56 | 61 | 0.46 |
| CytC | 5 | 69 | 65 | 0.61 | 88 | 0.63 |
| CytC | 10 | 69 | 65 | 2.27 | 88 | 0.29 |
| Collagen | 2.5 | 31 | 34 | 0.69 | 61 | 1.6 |
| Collagen | 5 | 31 | 32 | 1.26 | 61 | 3.04 |
| Collagen | 2.5 | 50 | 52 | 1.25 | 74 | 1.67 |
| Collagen | 5 | 50 | 52 | 1.03 | 75 | 0.91 |
| Collagen | 2.5 | 69 | 73 | 1.08 | 86 | 0.85 |
| Collagen | 5 | 69 | 70 | 0.38 | 86 | 0.48 |

When analysed using ESI, peptide mixtures spiked into BSA tryptic digests showed the most variability between concentrations. Peptide mixtures spiked into CytC digests showed good reproducibility across the two concentrations, with ~ 4 and 5% underestimation of product peptide Q levels at 50 and 70 % Q respectively. Interestingly, given that collagen is the largest of the three proteins tested, when spiked into the archaeological collagen tryptic digests, the measurements of percentage Q showed the best consistency across the two concentrations. The measured percentage Q in collagen was accurate within 4%. When the same spiked protein digests were analysed using MALDI there was a large overestimation of peptide Q levels across all three sample types.

# Discussion

This study highlights the importance of investigating the suitability of analytical methodology to the question being asked: Including the suitability of the proposed mass spectrometric techniques used to answer research questions. The investigation of deamidation using mass spectrometry has been wildly applied and reported heavily in the literature, despite this, we are still discovering and learning, not only about the mechanism itself, but also about how the procedures we use to extract and analyse deamidation can influence the resulting Q/E measurements. With papers published only recently detailing the effects of well-established extractions procedures.[16,25] When deciding which ionisation source is most suitable for the samples in question, it is important to investigate relative mass spectrometric responses. Here we investigated to what extent two product peptides (one containing Q and one containing E, but otherwise identical in sequence), ionise in unseparated mixtures, in various ratios. Stapels and Barofsky (2004)[26] analysed a large number of peptides using both ESI and MALDI and found that there were groups of peptides that were only observed when analysed using one or the other ionisation technique. They observed that properties such as the isoelectric point of the peptide made little difference in the peptides’ behaviour using the two ionisation methods.  Among the peptides analysed by Stapels and Barosfsky (2004)[26], peptides containing glutamic acid were found to have statistically significant differences in ionisation behaviour between ionisation sources and the number of times they were observed, with peptides containing glutamic acid being preferentially ionised by ESI. In addition it has also been reported that asparagine and aspartic acid-containing peptides ionise similarly under the conditions of ESI.[27,28] The product peptides analysed in this paper showed ionisation bias when analysed using MALDI, with product peptide E consistently underestimated. Analysis of the same mixtures however using ESI showed no ionisation bias between the two peptide products. When using ESI the full range of product peptide mixtures studied was measured accurately. This accuracy was maintained when the mixtures were spiked into protein digests of small (CytC), medium (BSA) and large (collagen) proteins, at different concentrations.

There are a few possible reasons for the observed ionisation bias. It could be the way that the peptides are incorporated into the matrix, or how they behave in the plume on desorption. When analysed using LC-MS, the two peptide forms did resolve chromatographically, however in a biological mixtures such as protein digests there may be coelution of other Q and E containing peptides. If possible, chromatographic separation of Q and E containing peptides prior to MS analysis could help to negate potential ionisation bias.

Although this study originated with questions raised over the wide adoption of MALDI rather than ESI for analysis of ancient protein samples, the results are of much broader applicability and contribute to the body of information on peptide ionisation bias that is relevant to all those working in peptide mass spectrometric analysis, since sample handling procedures can be a cause of deamidation (for example protocols used in preparation of samples for proteomic analysis, or during production of biotherapeutics). Demonstrating that ESI is much better able to generate reliable response ratios of deamidated to undeamidated peptides than MALDI is an important observation, reaching well beyond the ancient protein community.

# Conclusions

Two peptides, differing only in that one contains an internal Q residue, and the other an E at the same position, were synthesised in high purity using solid phase synthesis. The synthesis process was very successful and yielded peptide products that were significantly purer than many commercially available ‘authentic standards’.

It appears that, when choosing an ionisation source, ESI is more suitable than MALDI for accurately measuring ratios of Q- and E-containing peptides. We have shown that use of MALDI results in an under-representation of product peptide E across the full range of different percentage ratios. The level of underestimation of product peptide E varied across different peptide ratios, so that use of a simple correction factor is not sufficient to mitigate this. It should be noted that the effects of ionisation have only been studied here on one peptide sequence, and that the amino acid sequence may also have an effect on ionisation bias. However, it has been shown here that even using mixtures of synthetic products containing essentially just two peptides (Product peptides Q and E), preferential ionisation of product peptide Q occurs when using MALDI. This is observation is also valid in more complex peptide mixtures, such as a digest of a single protein. This observation alone is sufficient to question continued automatic use of MALDI in preference to ESI for such analyses.  Our preliminary observations make it clear that further investigation is required, not only to look at different peptide sequences, but also to investigate the suitability of ESI when looking at Q to E ratios in other complex biological mixtures such digests of protein mixtures as well as other biological matrices such as tissue samples used for imaging.

Acknowledgements

We would like to thank the following: NERC (NE/J500197/1) for providing the funding for this PhD research, Adam Dowle (Technology Facility, University of York) and Sheila Taylor (Department of Chemistry, University of York) for their continuing technical support and advice. We gratefully acknowledge use of both the ultraflex and solariX instruments in the York Centre of Excellence in Mass Spectrometry, which was created thanks to a major capital investment through Science City York, supported by Yorkshire Forward with funds from the Northern Way Initiative. Lastly, we would like to thank the Leverhulme Trust for additional support.

# References

[1] Robinson NE, Robinson A. Molecular clocks: deamidation of asparaginyl and glutaminyl residues in peptides and proteins. Althouse press. 2004.

[2] Li X, Lin C, O’Connor PB. Glutamine deamidation: differentiation of glutamic acid and γ-glutamic acid in peptides by electron capture dissociation. Analytical Chemistry. 2010;82 (9): 3606–15. doi: 10.1021/ac9028467

[3] Kuriakose A, Chirmule N, Nair P. Immunogenicity of Biotherapeutics: Causes and Association with Posttranslational Modifications. Journal of Immunology Research. 2016;20

doi: 10.1155/2016/1298473

[4] Takemoto L, Boyle D. Deamidation of Specific Glutamine Residues from Alpha-A Crystallin during Aging of the Human Lens. Biochemistry. 1998;(37) 13681–13685.

[5] Flaugh SL, Milk IA, King J. Glutamine Deamidation Destabilizes Human γD-Crystallin and Lowers the Kinetic Barrier to Unfolding. The Journal of Biological Chemistry. 2006;(281): 30782–30793

doi: 10.1074/jbc.M603882200

[6] Adav SS, Subbaiaih RS, Kerk, SK et al., Studies on the Proteome of Human Hair - Identification of Histones and Deamidated Keratins. Scientific Reports. 2018: 1-11. doi:10.1038/s41598-018-20041-9

[7] Serra A, Gallart-Palau X, Wei J et al.., Characterization of Glutamine Deamidation by Long-Length Electrostatic Repulsion-Hydrophilic Interaction Chromatography-Tandem Mass Spectrometry (LERLIC-MS/MS) in Shotgun Proteomics. Analytical Chemistry. 2016; 88 (21): 10573-10582. doi: 10.1021/acs.analchem.6b02688

[8] Huang L, Lu J, Wroblewski VJ, et al., In Vivo Deamidation Characterization of Monoclonal Antibody by LC/MS/MS. Analytical Chemistry. 2005;77 (5): 1432–1439. doi: 10.1021/ac0494174

[9] Chelius D, Rehder DS and Bondarenko PV. Identification and Characterization of Deamidation Sites in the Conserved Regions of Human Immunoglobulin Gamma Antibodies

Analytical chemistry. 2005;77 (2005): 6004-6011. doi.org/10.1021/ac050672d

[10] Zhang A, Hu P, MacGregor P, et al., Understanding the Conformational Impact of Chemical Modifications on Monoclonal Antibodies with Diverse Sequence Variation Using Hydrogen/Deuterium Exchange Mass Spectrometry and Structural Modeling. Analytical Chemistry. 2014; 86 (7): 3468-3475. doi: 10.1021/ac404130a

[11] Demarchi B, Hall S, Roncal-Herrero, et al. Protein sequences bound to mineral surfaces persist into deep time. Elife 2016; 5: e17092.2016. doi: 10.7554/eLife.17092.001

[12] Araki N and Moini, M. Age stimation of museum wool textiles from Ovis aries using deamidation rates utilizing matrix-assisted laser desorption/ionization time-of-flight mass spectrometry. Rapid communications in mass spectrometry. 2011;25 (22): 3396–3400.  doi: 10.1002/rcm.5237

[13] Solazzo C, Wilson J, Dyer JM, et al. Modeling Deamidation in Sheep α-Keratin Peptides and Application to Archeological Wool Textiles. Analytical Chemistry. 2014;86 (1): 567–575.                              doi: 10.1021/ac4026362

[14] Mikšík I, Sedláková P, Pataridid S, et al., Prince Cangrande’s Collagen: Study of Protein Modification on the Mummy of the Lord of Verona, Italy (1291–1329 AD). Chromatographia. 2014; (77) 1503–1510. doi: 10.1007/s10337-014-2710-0

[15] van Doorn NL, Wilson J, Hollund H, et al., Site-specific deamidation of glutamine: a new marker of bone collagen deterioration. Rapid Communications in Mass Spectrometry. 2012;(26) 2319–2327. doi: 10.1002/rcm.6351

[16] Simpson JP, Penkman KEH, Demarchi B, et al. The effects of demineralisation and sampling point variability on the measurement of glutamine deamidation in type I collagen extracted from bone. Journal of Archaeological Science. 2016;(69) 29-38. doi: 10.1016/j.jas.2016.02.002

[17] Leo G, Bonadice I, Andreotti A et al., Deamidation at asparagine and glutamine as a major modification upon deterioration/aging of proteinaceous binders in mural paintings. Analytical Chemistry. 2011;(83) 2056–2064. doi: 10.1021/ac1027275

[18] Joshi AB and Kirsch LE. The estimation of glutaminyl deamidation and aspartyl cleavage rates in glucagon. International journal of pharmaceutics. 2004;(273) 213-219.  doi:10.1016/j.ijpharm.2004.01.006

[19] Joshi AB, Sawai M, Kearney WR et al., Studies on the mechanism of acid cleavage and glutamine deamidation in the acidic degradation of glucagon. Journal of pharmaceutical sciences. 2005;94 (9): 1912-1927. doi: 10.1002/jps.20405

[20] Geiger T and Clarke S. Deamidation, isomerization, and racemization at asparaginyl and aspartyl residues in peptides. Succinimide-linked reactions that contribute to protein degradation. Journal of Biological Chemistry. 1987;(262) 785-794.

[21] Li X, Cournover JJ, Lin C, O’Connor PB. Use of 18O labels to monitor deamidation during protein and peptide sample processing. Journal of American Mass Spectrometry. 2018;19 (6): 855-64. doi: 10.1016/j.jasms.2008.02.011

[22] Nilsson MR, Driscoll M, Raleigh DP. Low levels of asparagine deamidation can have a dramatic effect on aggregation of amyloidogenic peptides: Implications for the study of amyloid formation. Protein Science. 2002;11(2): 342–349. doi: 10.1110/ps.48702

[23] Vanbelle C, Halgand F, Cedervall T, et al. Deamidation and disulfide bridge formation in human calbindin D28k with effects on calcium binding. Protein Science. 2005;14(4): 968–979. doi:10.1110/ps.041157705

[24] Wilson J, van Doorn NL, Collins MJ. Assessing the Extent of Bone Degradation Using Glutamine Deamidation in Collagen. Analytical Chemistry. 2012;(84) 9041–9048. doi: 10.1021/ac301333t

[25] von Holstein ICC, von Tersch M, Couto AN, et al., Collagen proteins exchange oxygen with demineralisation and gelatinisation reagents and also with atmospheric moisture. Rapid Communications in Mass Spectrometry. 2018; (32):523 –534.                    doi: 10.1002/rcm.8064

[26] Stapels MD, and Barofsky D. Complementary use of MALDI and ESI for the HPLC-MS/MS analysis of DNA-binding proteins. Analytical chemistry, 2014;76 (18): 5423-5430. doi: 10.1021/ac030427z

[27] Stroop S. A modified peptide mapping strategy for quantifying site-specific deamidation by electrospray time-of-flight mass spectrometry. Rapid Communications in Mass Spectrometry. 2007;(21) 830–836. doi: 10.1002/rcm.2901

[28] Hurtado PP and O’Connor PB. Deamidation of collagen. Analytical Chemistry. 2012;(84) 3017–3025. doi: 10.1021/ac202980z